## 単目量
製造事業者名： | [6 empty] | アナログ出力モジュールの適合性チェック －単目量はかり－
Type：
（1）ロードセル（LC），指示計（IND）及びはかり（WI）の精度等級
[8 empty] | 3 |  | 3 |  |  |  | 3
[8 empty] | LC | & | IND | 以上 |  |  | WI |  | 適合 | 不適合
非自動はかり (Weighing instrument) | 精度等級 （accuracy class) | class |  | ３級 |  | - |  | C | & | ３級 | 以上 |  |  | ３級 |  | ☑
| ひょう量 (maximum capacity) | Max |  | 60000 |  | kg | （2）ロードセル（LC）及び指示計（IND）の使用温度範囲並びに比較したはかり（WI）の使用温度範囲
| 目量 (verification scale interval) | e |  | 20 |  | kg |  |  | LC |  | IND |  |  | WI |  | 適合 | 不適合
| ロードセルの数 (number of load cells) | N |  | 4 |  | - |  | Tmin | -10 | & | -10 |  | ≦ | -10 | ℃ | ☑
| 減少率 (reduction ratio) | R |  | 1 |  | - |  | Tmax | 40 | & | 40 |  | ≧ | 35 | ℃ | ☑
| 荷重受け部の荷重 (dead load of load receptor) | 下限 (DLmin) | 上限 (DLmax) | 12000 | 15000 | kg | （3）接続要素，指示計及びロードセルの最大許容誤差の誤差配分piの平方根の和
| 初期零点設定範囲 (initial zero setting range) | IZSR |  | 10000 |  | kg |  | pcon2 | + | pind2 | + |  | pLC2 | ≦ 1 |  | 適合 | 不適合
| 不均一分布荷重補正 (correction for non uniform distributed load) | NUD |  | 30000 |  | kg |  | 0.25 | + | 0.25 | + |  | 0.49 | ≦ 1 |  | ☑
| 加算式風袋量 (additive tare) | T+ |  | 0 |  | kg | （4）指示計の目量の数の最大数及びはかりの目量の数
| 使用温度範囲 (temperature range) | 下限 (Tmin) | 上限 (Tmax) | -10 | 35 | ° C |  |  |  |  | nind |  | ≧ | ni=Maxi/ei |  | 適合 | 不適合
| ケーブルの長さ (length of connecting cable) | L |  | 100 |  | m |  |  | 単目量はかり |  | 3000 |  | ≧ | 3000 |  | ☑
| ケーブルの心線の断面積 (cross section of wire) | A |  | 0.75 |  | mm2 | （5）ロードセルの最大定格荷重とはかりのひょう量との関係
| IZSR：ひょう量の10%, NUD：トラックスケール＝ひょう量の50%、台はかり＝ひょう量の20% | [5 empty] | 係数Q: Q＝(Max＋DLmax＋IZSR＋NUD＋T＋)/Max＝ | [6 empty] | 1.9167 |  |  |  | 1.8667
指示装置 (Indicator) | 精度等級 (suitable for accuracy class of the weighing instrument) | class |  | ３級 |  | - |  |  |  | Q×Max×R/N |  |  | ≦ | Emax |  | 適合 | 不適合
| 目量数の最大数 (maximum number of verification scale interval) | nind |  | 3000 |  | - |  |  |  | 28750 |  | kg | ≦ | 30000 | kg | ☑
製造事業者名 | ロードセルの印加電圧 (load cell excitation voltage) | Uexc |  | 12 |  | V | （6a）ロードセルの目量の数の最大数及びはかりの目量の数
| 最小入力電圧 (minumum input voltage) | Umin |  | 1 |  | mV |  |  |  |  | nLC |  | ≧ | ni=Maxi/ei |  | 適合 | 不適合
型式 | 目量当たりの最小入力電圧 (minimum input voltage per verification scale interval) | Δumin |  | 1 |  | μV |  |  | 単目量はかり |  | 3000 |  | ≧ | 3000 |  | ☑
| 最小・最大ロードセルインピーダンス (load cell resistance) | 最小 (RLmin) | 最大 (RLmax) | 30 | 1000 | Ω | （6d）ロードセルの最小荷重及び荷重受け部（kg）
成績書番号 | 使用温度範囲 (temperature range) | 下限 (Tmin) | 上限 (Tmax) | -10 | 40 | ° C |  |  |  |  | DLmin×R/N |  | ≧ | Emin |  | 適合 | 不適合
| 誤差配分 (fraction of the maximum permissible error) | pind |  | 0.5 |  | - |  |  |  |  | 3000 | kg | ≧ | 2000 | kg | ☑
| ケーブル接続 (cable connection) |  |  | 6 |  | 線式 | （7）はかりの目量とロードセルの目量との関係
| ケーブルの心線の断面積当たりの長さ (maximum value of cable length per wire cross section) | (L/A)max |  | 150 |  | m/mm2 |  |  |  |  | e×R/√N |  | ≧ | vmin=Emax/Y |  | 適合 | 不適合
[11 empty] | 10 | kg | ≧ | 5 | kg | ☑
ロードセル (LC) (Load cell(s)) | 精度等級 (accuracy class) | Class |  | C |  | - | （8）指示計に対する一般的な最小入力電圧及び目量当たりの最小入力電圧及びロードセルの出力
| 最大定格荷重 (maximum capacity) | Emax |  | 30000 |  | kg |  |  | U =C×Uexc×R×DLmin/(Emax×N) |  |  |  | ≧ | Umin |  | 適合 | 不適合
製造事業者名 | 最小荷重 (minimum dead capacity) | Emin |  | 2000 |  | kg |  |  | 2.4 |  | mV |  | ≧ | 1 | mV | ☑
| 定格出力 (rated output) | C |  | 2 |  | mV/V |  |  | Δu =C×Uexc×R×e/(Emax×N) |  |  |  | ≧ | Δumin |  | 適合 | 不適合
型式 | 目量の数の最大数 (maximum number of verification scale intervals) | nLC |  | 3000 |  | - |  |  | 4 |  | μV |  | ≧ | 1 | μV | ☑
| ロードセルの最小目量 (minimum verification scale interval) 又は | vmin |  |  |  | kg | （9）指示計に対する許容インピーダンス範囲及び実ロードセルインピーダンス（Ω）
成績書番号 | Y = 最大定格荷重 ÷ 最小目量 (Y = Emax / Vmin) | Y |  | 6000 |  | - |  |  | RLmin | ≦ | RLC / N |  | ≦ | RLmax |  | 適合 | 不適合
| 最小荷重の出力の戻り (minimum dead load output return) 又は | DR |  |  |  | kg |  |  | 30 | ≦ | 87.5 |  | ≦ | 1000 |  | ☑
| 最小荷重の出力の戻りとの比 (ratio Z = E max /(2 DR)) | Z |  | 3000 |  | - | (10) ケーブルの芯断面当たりのロードセルと表示器間の延長ケーブル長(m/mm²)
| ロードセルの入力抵抗 (input resistance of single load cell) | RLC |  | 350 |  | Ω |  |  |  |  | (L/A) |  | ≦ | (L/A)max |  | 適合 | 不適合
| 使用温度範囲 (temperature range) | 下限 (Tmin) | 上限 (Tmax) | -10 | 40 | ° C |  |  |  |  | 133.3333 |  | ≦ | 150 |  | ☑
| 誤差配分 (fraction of the maximum permissible error) | pLC |  | 0.7 |  | -
接続素子、その他 | 誤差配分 (fraction of the maximum permissible error) | pcon |  | 0.5 |  | -

## Two ranges WI (hidden)
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - Two ranges weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (IND) and load cell (LC) | [6 empty] | I. O. ?
| 3 |  | 0 |  |  | 0
| LC | & | IND | equal or better |  | WI
| C | & | ３級 | equal or better |  | ３級 | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin | -10 | & | -10 | £ | -10 | YES
| Tmax | 40 | & | 40 | ³ | 35 | YES
(3)Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
| 0.25 | + | 0.25 | + | 0.49 | £ 1 | YES
(4)Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n(i)=Max(i)/e(i)
| Einbereichswaage |  |  | 3000 | ³ | #REF!
| two ranges weighing instrument |  | i = 1 | 3000 | ³ | #REF! | #REF!
|  |  | i = 2 | 3000 | ³ | #REF! | #REF!
|  |  | i = 3 | 0 | ³ | 0
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument | [7 empty] | Hilfsfeld
Faktor Q (EN 45 501 No 4.12.1) : Q = (Max2+DL+IZSR+NUD+T+)/Max2 = | [5 empty] | #REF! |  | #REF!
|  |  | (Q*Max2*R)/N |  | £ | Emax
|  |  | #REF! |  | £ | 30000 | #REF!
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n(i)=Max(i)/e(i)
| Einbereichswaage |  |  | 3000 | ³ | #REF!
| two ranges weighing instrument |  | i = 1 | 3000 | ³ | #REF! | #REF!
|  |  | i = 2 | 3000 | ³ | #REF! | #REF!
|  |  | i = 3 | 0 | ³ | 0
(6b) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrteilungswaage
|  |  | n LC oder Z=Emax/(2*DR) |  | ³ | Maxr / e1
|  |  | 0 |  | ³ | -0.28571
(6c) Minimum dead load output return of the load cell and smallest verification scale interval e1 of a multiple range weighing instrument
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | 0,4*Max2/e1
|  |  | 3000 |  | ³ | #REF! | #REF!
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
|  |  |  | 3000 | ³ | 2000 | YES
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e1*R/ÖN | ³ | vmin= Emax/Y
|  |  |  | #VALUE! | ³ | 5 | #VALUE!
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U= C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) | 2.4 |  | ³ | 1 | YES
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
|  |  | #VALUE! |  | ³ | 1 | #VALUE!
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 30 | £ | 87.5 | £ | 1000 | YES
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
|  |  |  | 133.3333 | £ | 150 | YES
Eichanweisung 9, Abschnitt 9.1 |  |  |  | 2001-12-13 |  |  | Seite 3 von 6
Evaluation Two ranges weighing instrument

## Three ranges WI (hidden)
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - Three ranges weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (IND) and load cell (LC) | [6 empty] | I. O. ?
| 3 |  | 0 |  |  | 0
| LC | & | IND | equal or better |  | WI
| C | & | ３級 | equal or better |  | ３級 | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin | -10 | & | -10 | £ | -10 | YES
| Tmax | 40 | & | 40 | ³ | 35 | YES
(3)Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
| 0.25 | + | 0.25 | + | 0.49 | £ 1 | YES
(4)Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n(i)=Max(i)/e(i)
| Mehrbereichswaage |  |  | 3000 | ³ | #REF!
| three ranges weighing instrument |  | i = 1 | 3000 | ³ | #REF! | #REF!
|  |  | i = 2 | 3000 | ³ | #REF! | #REF!
|  |  | i = 3 | 3000 | ³ | #REF! | #REF!
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument | [7 empty] | Hilfsfeld
Last-Korrekturfaktor: Q = (Max3+DL+IZSR+NUD+T+)/Max3 = | [5 empty] | #REF! |  | #REF!
|  |  | (Q*Max3*R)/N |  | £ | Emax
|  |  | #REF! |  | £ | 30000 | #REF!
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n(i)=Max(i)/e(i)
| Einbereichswaage |  |  | 3000 | ³ | #REF!
| three ranges weighing instrument |  | i = 1 | 3000 | ³ | #REF! | #REF!
|  |  | i = 2 | 3000 | ³ | #REF! | #REF!
|  |  | i = 3 | 3000 | ³ | #REF! | #REF!
(6b) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrteilungswaage
|  |  | n LC oder Z=Emax/(2*DR) |  | ³ | Maxr / e1
|  |  | 0 |  | ³ | -0.28571
(6c) Minimum dead load output return of the load cell and smallest verification scale interval e1 of a multiple range weighing instrument
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | 0,4*Max3/e1
|  |  | 3000 |  | ³ | #REF! | #REF!
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
|  |  |  | 3000 | ³ | 2000 | YES
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e1*R/ÖN | ³ | vmin= Emax/Y
|  |  |  | #VALUE! | ³ | 5 | #VALUE!
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U= C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) | 2.4 |  | ³ | 1 | YES
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
|  |  | #VALUE! |  | ³ | 1 | #VALUE!
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 30 | £ | 87.5 | £ | 1000 | YES
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
|  |  |  | 133.3333 | £ | 150 | YES
Eichanweisung 9, Abschnitt 9.1 |  |  |  | 2001-12-13 |  |  | Seite 4 von 6
Evaluation Three ranges weighing instrument

## Two-interval WI (hidden)
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - Two-interval weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (IND) and load cell (LC) | [6 empty] | I. O. ?
| 3 |  | 0 |  |  | 0
| LC | & | IND | equal or better |  | WI
| C | & | ３級 | equal or better |  | ３級 | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin | -10 | & | -10 | £ | -10 | YES
| Tmax | 40 | & | 40 | ³ | 35 | YES
(3)Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
| 0.25 | + | 0.25 | + | 0.49 | £ 1 | YES
(4)Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n(i)=Max(i)/e(i)
| Einbereichswaage |  |  | 3000 | ³ | #REF!
| two-interval weighing instrument |  | i = 1 | 3000 | ³ | #REF! | #REF!
|  |  | i = 2 | 3000 | ³ | #REF! | #REF!
|  |  | i = 3 | 0 | ³ | 0
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument | [7 empty] | Hilfsfeld
Faktor Q (EN 45 501 No 4.12.1) : Q = (Max2+DL+IZSR+NUD+T+)/Max2 = | [5 empty] | #REF! |  | #REF!
|  |  | (Q*Max2*R)/N |  | £ | Emax
|  |  | #REF! |  | £ | 30000 | #REF!
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n(i)=Max(i)/e(i)
| Einbereichswaage |  |  | 3000 | ³ | #REF!
| two-interval weighing instrument |  | i = 1 | 3000 | ³ | #REF! | #REF!
|  |  | i = 2 | 3000 | ³ | #REF! | #REF!
|  |  | i = 3 | 0 | ³ | 0
(6b) Minimum dead load output return of the load cell and smallest verification scale interval e1 of a multi-interval weighing instrument
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | Max2 / e1
|  |  | 3000 |  | ³ | #REF! | #REF!
(6b) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrbereichswaage
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | 0,4*Maxr/e1
|  |  | 0 |  | ³ | 0
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
|  |  |  | 3000 | ³ | 2000 | YES
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e1*R/ÖN | ³ | vmin= Emax/Y
|  |  |  | #VALUE! | ³ | 5 | #VALUE!
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U= C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) | 2.4 |  | ³ | 1 | YES
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
|  |  | #VALUE! |  | ³ | 1 | #VALUE!
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 30 | £ | 87.5 | £ | 1000 | YES
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
|  |  |  | 133.3333 | £ | 150 | YES
Eichanweisung 9, Abschnitt 9.1 |  |  |  | 2001-12-13 |  |  | Seite 5 von 6
Evaluation Two-interval weighing instrument

## Three-interval WI (hidden)
Eichbehörden der Länder, englischsprachige Fassung
b) Verification of Compatibility - Three-interval weighing instrument (WI)
(1) Accuracy class of weighing instrument (WI) compatible to class of indicator (IND) and load cell (LC) | [6 empty] | I. O. ?
| 3 |  | 0 |  |  | 0
| LC | & | IND | equal or better |  | WI
| C | & | ３級 | equal or better |  | ３級 | NO
(2) Temp.limits of the weighing instr.(WI) compared with the temp.limits of the load cell (LC) and the indicator (IND)
|  | LC |  | IND |  | WI
| Tmin | -10 | & | -10 | £ | -10 | YES
| Tmax | 40 | & | 40 | ³ | 35 | YES
(3)Sum of the squares of the fractions pi of the max. permissible errors of connecting elements, indicator and load cells
| pcon2 | + | pind2 | + | pLC2 | £ 1
| 0.25 | + | 0.25 | + | 0.49 | £ 1 | YES
(4)Number of verification scale intervals of the weighing instrument and the indicator
|  |  |  | nind | ³ | n(i)=Max(i)/e(i)
| Einbereichswaage |  |  | 3000 | ³ | #REF!
| three-interval weighing instrument |  | i = 1 | 3000 | ³ | #REF! | #REF!
|  |  | i = 2 | 3000 | ³ | #REF! | #REF!
|  |  | i = 3 | 3000 | ³ | #REF! | #REF!
(5) Maximum capacity of load cells must be compatible to Max of the weighing instrument | [7 empty] | Hilfsfeld
Faktor Q (EN 45 501 No 4.12.1) : Q = (Max3+DL+IZSR+NUD+T+)/Max3 = | [5 empty] | #REF! |  | #REF!
|  |  | (Q*Max3*R)/N |  | £ | Emax
|  |  | #REF! |  | £ | 30000 | #REF!
(6a) Maximum number of verification scale intervals of load cell and number of scale intervals of the weighing instrument
|  |  |  | nLC | ³ | n(i)=Max(i)/e(i)
| Einbereichswaage |  |  | 3000 | ³ | #REF!
| three-interval weighing instrument |  | i = 1 | 3000 | ³ | #REF! | #REF!
|  |  | i = 2 | 3000 | ³ | #REF! | #REF!
|  |  | i = 3 | 3000 | ³ | #REF! | #REF!
(6b) Minimum dead load output return of the load cell and smallest verification scale interval e1 of a multi-interval weighing instrument
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | Max3 / e1
|  |  | 3000 |  | ³ | #REF! | #REF!
(6c) Rückkehr des Vorlastsignals der WZ und kleinster Eichwert e1 einer Mehrbereichswaage
|  |  | nLC oder Z=Emax/(2*DR) |  | ³ | 0,4*Maxr/e1
|  |  | 0 |  | ³ | 0
(6d) Minimum dead load of the load cells to the actual dead load of the load receptor
|  |  |  | DL*R/N | ³ | Emin
|  |  |  | 3000 | ³ | 2000 | YES
(7) Verification scale interval of the weighing instrument and minimum load cell scale interval must be compatible
|  |  |  | e1*R/ÖN | ³ | vmin= Emax/Y
|  |  |  | #VALUE! | ³ | 5 | #VALUE!
(8) Minimum input voltage for the indicator, minimum input voltage per verification scale interval and actual output of the LCs
|  | minimum input voltage | U= C*Uexc*R*DL/(Emax *N) |  | ³ | Umin
|  | (unloaded WI) | 2.4 |  | ³ | 1 | YES
|  | input voltage per verification interval | Du = C*Uexc*R*e/(Emax *N) |  | ³ | Dumin
|  |  | #VALUE! |  | ³ | 1 | #VALUE!
(9) Allowed impedance range for the electronic indicator and actual load cell impedance
|  | RLmin | £ | RLC / N | £ | RLmax
|  | 30 | £ | 87.5 | £ | 1000 | YES
(10) Cable length per wire cross section of the connection cable between the load cell(s) and indicator
|  |  |  | (L/A) | £ | (L/A)max
|  |  |  | 133.3333 | £ | 150 | YES
Eichanweisung 9, Abschnitt 9.1 |  |  |  | 2001-12-13 |  |  | Seite 6 von 6
Evaluation Three-interval weighing instrument

## Tabelle1 (hidden)
One range WI
Two ranges WI
Three ranges WI
Two-interval WI
Three -interval WI
1